## Sheet1
S No | Complaint Text | Resolution | Is Investigation Reqd | Is it required to report to Authorities
1 | A leakage was observed when using Alinity i-series Concentrated Wash Buffer. | A malfunction was identified regarding a leakage in Alinity i-series Concentrated Wash Buffer. Based on the information within the complaint record, the device failed to meet performance specifications.However, a systemic issue and/or product deficiency was not identified | No | No
2 | The customer indicated product shipment was received warm (temperature stressed) and therefore not suitable for use | A malfunction was identified regarding ARCHITECT Concentrated Wash Buffer was damaged in transit by FedEx since the shipment was received warm (temperature stressed) and therefore not suitable for use. No additional issues were identified; no further action is required at this time. | No | No
3 | Shipment of ARCHITECT Concentrated Wash Buffer, list number 06C54-88, lot number 08381FN00, was damaged in transit by FedEx and refused by the customer | A malfunction regarding Shipment of ARCHITECT Concentrated Wash Buffer was identified.The device failed to meet performance specifications or otherwise perform as intended at the customer site.However, a systemic issue and/or product deficiency was not identified. | No | No
4 | Alinity ci series System Software v2.6.2 on SCM01493 where the westgard rules aren’t triggered while there are two versions of the same assay installed on the instrument. | A malfunction was identified for the Alinity ci System Software 2.6.2.the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency of the software was not identified. PRS ICQ100391 is under investigation | Yes | No
5 | The customer observed bottles which appeared to be distended when using ARCHITECT Concentrated Wash Buffer Kit | No product deficiency and no malfunction was identified regarding distended while using ARCHITECT Concentrated Wash Buffer Kit . Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. Further investigation will be performed as per management discretion through pNCR706369. | Yes | No
6 | The customer observed activated read failures (error code 1007), when using ARCHITECT Reaction Vessels (RV) | A malfunction and a product deficiency have been identified regarding activated read failures while using ARCHITECT Reaction Vessels (RV).Quality Record (QR) has been opened to document a deficiency and further investigate this issue due to increased complaint activity being identified for the complaint issue | Yes | No
7 | The customer observed a Microparticles bottle containing ''mold-type deposits'', while using ARCHITECT Intact PTH reagent | A malfunction was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. The complaint text provided sufficient evidence that customer observed a Microparticles bottle containing ''mold-type deposits''.A deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the 'mold-type deposits'', while using ARCHITECT Intact PTH reagent. No additional issues were identified; no further action is required at this time. | No | No
8 | The customer reported observing falsely elevated LDL and HDL results for 5 patients when using Direct LDL reagent (LN 1E31-20) , Ultra HDL reagent (LN 3K33-21) on ARCHITECT 1601094 system. The sum of the LDL and the HDL results for each patient were greater than the total cholesterol results | Deficiency regarding falsely eleveted LDL and HDL could not be determined as the trend review and lot search did not identify an increase in complaint activity for the current issue. In addition, the customer’s QC results were within expected ranges. | No | No
9 | The customer observed multiple levels of non-Abbott controls shifting high while using ARCHITECT BNP calibrator | A malfunction and a product deficiency have been identified for the multiple levels of non-Abbott controls shifting high while using ARCHITECT BNP calibrator. The Quality Record (QR) has been opened to document this deficiency and further investigate this issue | Yes | No
10 | Assay package insert was not downloaded to the system when the Abbott Mail inbox was configured to receive inserts in the Norwegian language. The text states an assay insert in the English language was downloaded to the system when the Abbott Mail inbox language was reconfigured for English. The issue occurred on ALINITY system SCM02557 at v2.6.2 software | A malfunction was identified where the software failed to meet performance specifications or otherwise perform as intended at the customer site. However, there was insufficient information to indicate a deficiency of the ALINITY System Software v2.6.2. CLS Defect Tracking PRS CLS00435 will further document the issue | No | No
11 | The customer observed when they received the Troponin controls Lot: 89216UI00 that the bottle appeared misshapen, once the lid was removed the bottle returned to its normal shape | A malfunction was identified regarding misshapen of the bottle. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
12 | The complaint text stated that a Sodium result for sample ID 0425023014 was associated with sample ID 042302301. The result was not reported to the LIS because it was blocked by a rule in AlinIQ AMS which evaluated the Delta-Check of two reruns of the same sample ID | It is neither a malfunction nor a deficiency of the AlinIQ AMS LabOnline middleware was identified. The AlinIQ AMS middleware was performing as designed, and the issue was due to a customer IT network issue. | No | No
13 | The customer observed falsely elevated Vitamin D results when using the Alinity i Vitamin D reagent Lot: 96200UI00. The customer stated initial results were flagged > outside of linearity and repeated in the normal range. There was no impact to patient management reported. The issue was resolved with a new reagent Lot: 95489UI00 | A malfunction was identified regarding falsely elevated Vitamin D while using the Alinity i Vitamin D reagent Lot 96200UI00 failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
14 | The complaint text states that the when the customer adds a test to a sample that has an automatic rerun, the instrument does not add it, instead the request remains in pending. The customer is on ALINITY System Software v2.6.2. The instruments logs were not provided, and a customer database was not available; therefore, there is insufficient information to further evaluate the issue. | Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. The ALINITY System Software requirements indicate that the software is performing as designed. Neither a malfunction nor a deficiency of the ALINITY System Software v2.6.2 was identified | No | No
15 | Customer received a bent/damaged PROBE NEXT R from an unopened box. This is considered a first use failure of the PROBE NEXT R (list number 03R85-01) and is the subject of this evaluation. The part was replaced to resolve the issue. | A malfunction was identified regarding bent/damaged PROBE NEXT R from an unopened box. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site as the PROBE NEXT R had to be replaced to resolve the issue. No systemic issue or product deficiency was identified for the PROBE NEXT R | No | No
16 | The customer observed process path jams and/or instrument step loss errors, when using ARCHITECT RVs | A malfunction and a product deficiency have been identified for process path jams and/or instrument step loss errors when using ARCHITECT RVs. The Quality Record (QR) number 678997 has been opened to document this deficiency and further investigate this issue | Yes | No
17 | Auto dilution results are being transmitted first to the LIS instead of the original/neat results. The customer is on Alinity ci System Software v2.6.2 | A malfunction was identified for the auto dilution results are transmitted first to LIS instead of original/neat result in Alinity ci System Software v2.6.2. However, a deficiency of the software was not identified. PRS ICQ009359, further documents the issue, and the resolution will be addressed in the next software release. | No | No
18 | The customer observed low or imprecise/ erratic results or calibration issues when testing with the ARCHITECT Free T3 assay | A product deficiency has been identified for low or imprecise/ erratic results or calibration issues when testing with the ARCHITECT Free T3 assay.The root cause for this issue is under investigation. Further corrective actions will be implemented and will be communicated upon completion of the investigation. | Yes | No
19 | Customer was performing 6070 Daily maintenance and obtained error code 9034 Call Abbot, Sampler Handler status does not allow execution of command and error code 9071 Run-time error encountered while running (6070 Daily Maintenance) procedure. Customer was able to complete Daily Maintenance after the 4th attempt and instrument is up and running. The issue occurred on C401537 at software version 8.1. | A malfunction was identified with the Architect System Software v8.10. However, there is insufficient information to indicate a deficiency with the Architect System Software v8.10. | No | No
20 | The customer stated the APS Reseal Aluminum Coil (List Number 07L12-51), lot number RN-18-05608, would not seal and appeared to be wound incorrectly within the cartridge. This resulted in a first use failure of the part. Replacement of the coil resolved the issue. A photo was provided to show evidence of the failure. | Investigation of the issue through complaint identified a deficiency as an adverse trend was found for all lot numbers involving the resealer coil not advancing properly. The current complaint describes a similar failure for the coil being too wide. | No | No
21 | Deformed microparticle bottle observed when using ARCHITECT HIV Ag/Ab Combo reagent, lot number 04485BE03 | A malfunction was identified due to a deformed microparticle bottle was observed. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
22 | The customer observed that the Quality Controls at Levels 1 are out of range low when using ARCHITECT TSH Reagent | A malfunction of the ARCHITECT TSH Reagent was identified. Based on the information provided, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
23 | The customer observed microparticle clumping/resuspension issue for ARCHITECT Anti-HBs | A malfunction was identified regarding microparticle clumping/resuspension. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | Yes | No
24 | Negative Control too high due to aberrant calibration when using Alinity i Toxo IgG Calibrator Kit, list number 7P45-01, lot number 94114LI00 | A malfunction was identified regarding high negative control. A systemic issue and/or product deficiency was not identified.The customer observed a too high Negative Control (reactive result) after calibration with lot number 94114LI00; following calibration with calibrator lot number 03176BE00 gave an acceptable result for the negative control (0.0 IU/mL for 3 replicates). For the implicated calibration, the customer observed an aberrantly high replicate for Calibrator A. | No | No
25 | The complaint text states that a calibration request for sodium, potassium and chloride remains ''in process'' in the calibration status screen, even though QC is run and validated. The customer resolves the ''in process'' status, by cycling power to the instrument and then recalibrates the sodium, potassium and chloride. The customer is on ALINITY System Software v2.6.1. | A malfunction was identified where the software failed to meet performance specifications or otherwise perform as intended at the customer site. However, there was insufficient information to indicate a deficiency of the ALINITY System Software v2.6.1 | No | No
26 | Customer observed low results in there BioRad Controls when using ARCHITECT FOLATE Reagent (Lot: 87444UI00), resulting in out of range low Quality Controls at level 1. The customer recalibrated with a New Reagent Lot and the issue was resolved. | A malfunction of the ARCHITECT FOLATE Reagent was identified. Based on the information provided, the device failed to meet performance specifications or otherwise perform as intended at the customer site.However, a systemic issue and/or product deficiency was not identified. | No | No
27 | The customer called to report that they ordered 4 boxes of Ethanol Reagent LN 03L36-20, Lot 60099FD01 and 2 boxes had empty R1 wedges. Customer requesting replacement of the product | A malfunction was identified regarding 2 empty box of ethanol out of 4 boxes of ethanol ordered. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. No additional issues were identified; no further action is required at this time. | No | No
28 | The customer complains they are unable to load the calibrator values from the USB Card when using MCC Lot 83787FD01, List Number: 1E65. The customer tried multiple times without success. | A malfunction was identified regarding they are unable to load the calibrator values from the USB card. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.A Deficiency is identified. The USB Cal Value Media card provided in MCC Lot 83787FD01 is not useable by the customer. Refer to QR 685557 | Yes | No
29 | The customer observed non-Abbott controls and patient results shifting high while using ARCHITECT BNP calibrator lots 44K80818 , 44K82018 , 44K81018 and 44K85219 | A malfunction and a product deficiency have been identified for the complaint issue. The Quality Record (QR) number 678641 has been opened to document this deficiency and further investigate this issue. The likely cause lot number is listed as an impacted lot number in QR 678641. | Yes | No
30 | Customer states one bottle MCC Calibrator LN 1E65-05, Lot 29380FD01 is missing a label | A malfunction was identified regarding missing label in MMC Calibrator. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. No further action is required at this time | No | No
31 | The customer could not load the APS Reseal Aluminum Coil (List Number 07L12-51) on their Resealer Module. The coil was received without the protective seal and the foil was inside the cartridge, resulting in a first use failure of the part. Replacement of the coil resolved the issue. | Investigation of the issue through complaint ticket 225225RA1008301 identified a deficiency as an adverse trend was found for lot numbers RN-17-04939 and RN-17-04940. Lot number RN-18-05606 reported in the current complaint describe a similar failure to those in previous lot numbers RN-17-04939 and RN-17-04940. The failure prevents the Resealer Module from operating as described in the ACCELERATOR Automated Processing Systems Operations Manual (580799014.0, January 2013). The APS Reseal Aluminum Coil is part of the Resealer Module which is used on the APS, a3600 and the Inpeco FlexLab systems. Quality Record (QR) 642017 was initiated in the Abbott TrackWise for Quality (ABTRAQ) to track further activities for this issue; the scope of QR 642017 was expanded to include all lot numbers | Yes | No
32 | The complaint text states the software shutdown unexpectedly. The issue was resolved by rebooting the system. The issue occurred on Alinity SCM module SCM01207 operating at Alinity ci series System Software version 2.6.1. | A deficiency of the Alinity ci-series System Software v2.6.1 was not identified as the investigation determined there was no software requirement failure. A malfunction was identified as the system failed to meet performance specifications or otherwise perform as intended at the customer site. PRS ICQ008392 states the issue has been addressed in the Alinity ci software v2.6.2 release. Alinity ci software v2.6.2 was released for implementation in the field by TSB 214-028 on 07/02/2019. | No | No
33 | The complaint text states that the for multiple assays the customer has CNTL flags on results. Customer has deleted all control failures on all assays, but the CNTL flags are still present. The issue occurred on SCM01723 at Alinity ci series System Software v2.6.0. | A malfunction was identified for the Alinity System Software 2.6.0. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. Similar issues identified PRS ICQ009050. PRS ICQ009050 was opened to investigate why a CNTL flag cannot be cleared if the assay and control is uninstalled then reinstalled while in an active QC Failure state. PRS ICQ009050 indicates the issue is currently under investigation and the issue does not fail to meet a software requirement | Yes | No
34 | The customer observed the error code: 3000 ‘Unable to process test, Liquid not found for (Pipettor) at (R1 Middle Reagent)’ when using the ARCHITECT ESTRADIOL reagent Lot 95176UI00. The customer stated that they observed the error after 21 tests and after inspection they observed that the reagent bottle was empty. | A malfunction was identified regarding unable to process test because of the empty reagent bottle. Based on the information provided, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified | No | No
35 | The customer reported a concern regarding the Alinity analyzer, Serial number HQ00144, generating a 0107 Message Code ''Rack in lane (0) removed manually by the operator'' when racks had not been removed from the instrument | Based on the investigation, a product deficiency was identified for the Alinity hq, 09P68-01, because it has generated false Message Codes. A malfunction was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A Quality Report (QR) 650454 was opened to further evaluate the identified product issue. | Yes | No
36 | The customer observed activated read failures (error codes 1006 & 1007), when using ARCHITECT Reaction Vessels (RV) lot 000186937 | A malfunction and a product deficiency have been identified for the complaint regarding activated read failures.Quality Record (QR) number 698922 has been opened to document a deficiency and further investigate this issue due to increased complaint activity being identified for the complaint issue. | Yes | No
37 | The complaint text states that the ARCHITECT analyzer generated a falsely depressed CA 19-9XR result for a patient who had not been diagnosed with pancreatic cancer at the time of testing. | A cross functional team (CFT) was formed (as described in PPM ADD-CAPA-030) to review the product labeling and complaint activity. The CFT confirmed that the use described in this ticket was not in alignment with our labeling and that the current product labeling was reviewed and is sufficient and no further action will be taken at this time. Therefore, a deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
38 | The customer observed a falsely elevated result when using ARCHITECT Reaction Vessels (RV) lot 000195140. | A malfunction and a product deficiency regarding falsely elevated result when using ARCHITECT Reaction Vessels (RV) have been identified.Quality Record (QR) number 698922 has been opened to document a deficiency and further investigate this issue due to increased complaint activity being identified for the complaint issue. | Yes | No
39 | The complaint text states that while loading controls using the Pen drive the system software unexpectedly shut down. Customer cycled power to the SCM and modules to resolve the issue. The issue occurred on Alinity SCM02578 at software version 2.6.0 | A malfunction was identified with the Alinity ci System Software v2.6.0. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, there is insufficient information to indicate a deficiency with the Alinity ci System Software v2.6.0. Per PRS ICQ008392, the issue was addressed in ALINITY ci series System Software v2.6.2. | No | No
40 | The complaint text states that an ARCHITECT Concentrated Wash Buffer, lot 04275FN00, list 06C54-82 was damaged in transit. | A malfunction was identified regarding ARCHITECT concentrated Wash Buffer was damaged in transit. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
41 | The customer observed a control level 1 bottle that broke as soon as it was taken out to thaw, when using Technopath Multichem S Plus (Assayed) control lot 0178101701. | An investigation was performed by the Third Party Manufacturer (TPM), Technopath. The investigation by the TPM encompassed a review of batch records, and all paperwork relating to the manufacturing process for the likely cause control lot and found no issues. Retained vials were also inspected and found no issues. Furthermore, a review of batch records, CAPA and non-conformances were reviewed and did not identify any issues. As a result, this indicates that the product is performing as intended. A deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
42 | The customer observed Error code 3700, Unable to process test, (Wash Zone Aspiration 2) wash aspiration error for probe(s) (1) on ARCHITECT i2000SR, Serial Number ISR62638. It was determined that the ARC WZ Temperature Tubing/Sensor (List Number 08C94-90, Serial Number 2018) as Likely Cause | An adverse trend associated with Error 3700 was identified for the ARC WZ Temperature Tubing/Sensor, List Number 08C94-90. Potential Nonconformance QR 626346 was initiated on 03/01/2018 to further investigate the issue. A deficiency regarding the Error code 3700 was not identified. | Yes | No
43 | The customer observed microparticles that would not re-suspend, when using ARCHITECT Homocysteine reagent lot 10693UP00 | A deficiency regarding non re-suspend of microparticles while using ARCHITECT Homocysteine reagent was not identified as the trend review and lot search determined that complaint activity was normal for the complaint issue. No additional issues were identified; no further action is required at this time. | No | No
44 | Customer complains of higher than expected Calcium results for Abbott Controls with Reagent LN 3L79, Lots 31424UN18, 01781UN18 and 31423UN18 on their ARCHITECT Analyzer. | A deficiency reagrding higher Calcium result was not identified as testing shows recovery within expected ranges, the assay performs per specification. | No | No
45 | This complaint is for a calibration error associated with Rheumatoid Factor (RF) list number 06K44-01, lot number 01518H000. | A malfunction was identified regarding calibration error associated with Rheumatoid Factor (RF). Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. While the third party manufacturer determined no deficiency, the trend review and lot search identified an increase in complaint activity for the current issue. Per ADD-CAPA-038, QR 665920 was opened due to an Adverse complaint Trend (AssayTT-12217) to further investigate this issue and it will be considered a deficiency in our quality system. | Yes | No
46 | The customer indicated product shipment was received damaged and therefore not suitable for use. | A malfunction was identified , since the shipment was received damaged and therefore not suitable for use. The trend review identified normal complaint activity for the complaint issue and did not identify a systemic issue related to the current complaint, and therefore, a deficiency was not identified. No additional issues were identified; no further action is required | No | No
47 | The customer raised a complaint for ALINITY Probe conditioner, List Number 1R58-40, lot number 96338FN00 for a fill volume issue. In this case the customer observed message code #3424: Pipettor aspiration error. | A malfunction was identified regarding Pipettor aspiration error. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
48 | Positive Control shifted down when using ARCHITECT ANti-HCV reagent lot number 94361L100. The HCV Positive control is within range but showing a low trend. | A malfunction regarding Positive Control shifted down while using ANti-HCV reagent was not identified. Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. The customer observed shifted down values for the Positive Control when using the complaint lot but the values were within specifications.Based on our investigation we concluded that the ARCHITECT Anti-HCV reagent lot numbers 95367LI00/ 95371LI00 and 94357LI00/ 94361LI00 meet the performance specifications for the controls and sensitivity. | No | No
49 | The customer observed control level 3 bottles that were cloudy and contains particulates, when using Technopath Multichem IA Plus control lot 37104170. | A deficiency was not identified as the trend review and lot search determined that complaint activity was normal for the complaint issue. Additionally, an investigation was performed by the Third Party Manufacturer (TPM), Technopath. The investigation by the TPM encompassed Lot Confirmation testing and Internal Quality Control with no particulates observed in any of the vials run. In addition, a review of batch records and all paperwork relating to the manufacturing process were reviewed and did not identify any issues. Furthermore, a review of CAPA and non-conformances were reviewed and did not identify any issues. As a result, this indicates that the product is performing as intended. No additional issues were identified; no further action is required at this time. | No | No
50 | The customer reported that the conjugate bottle was empty upon first use when using a kit of ARCHITECT HBsAg Qualitative II, lot 87213FN00. The customers instrument returned error code 3000 (Unable to process test, Liquid not found for (x) at (y)) | A malfunction regarding conjugate bottle was empty upon first use when using a kit of ARCHITECT HBsAg Qualitative II was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
51 | The customer stated they observed a shift up in control results for multiple levels of Technopath Multichem IA+ lot 032308180 when using the Alinity i Folate reagent Lot: 04028UI00. A recalibration was completed, Technopath Multichem IA+ control values were Out of Range High after calibration. Alinity i Folate controls were testes and returned values within specification. | A malfunction regarding a shift up in control result was identified. Based on the information within the complaint record, the Alinity i Folate reagent Lot 04028UI00 failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | Yes | No
52 | Viscous matter visible within the conjugate bottle when using ARCHITECT Anti-HBe, lot number 86321LI00 | A malfunction was identified due to a viscous matter was visible within the conjugate bottle. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. |  | No
53 | The customer observed Technopath Control level 3 out of package insert range low with the ARCHITECT Direct Bilirubin assay, when using Technopath Multichem S Plus (Assayed) control lot 17205173. | A malfunction was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. The investigation by the TPM encompassed a review of batch records, quality control release data, stability data, and peer data for the ARCHITECT Direct Bilirubin assay and found no issues. Furthermore, a review of CAPA and non-conformances did not identify any issues. As a result, this indicates that the products are performing as intended. A deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
54 | The customer observed falsely elevated patient results that did not align with results obtained using other platform, while using ARCHITECT Anti-Tg reagent lot 67380UN18 . | A malfunction regarding falsely eleveted patient result was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.A deficiency was not identified as the trend review and lot search determined that complaint activity was normal for the customer's issue. No additional issues were identified; no further action is required at this time. | No | No
55 | The customer called to report their observation discrepant Magnesium result on the ARCHITECT c4000 Analyzer. The customer was running Clinical Chemistry Magnesium reagent LN 07D70-31 Lot 99662UN18. The complaint text listed the 12 falsely elevated results and confirmed that there was no impact to patient management. | There was no malfunction identified regarding discrepant Magnesium result on the ARCHITECT c4000 Analyzer. Based on the information within the complaint record, the device meets performance specifications and is performing as intended at the customer site.A systemic issue and/or product deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the current issue for the complaint lot. No additional issues were identified; no further action is required. |  | No
56 | The complaint text states that the customer loaded on a new lot of MCC (50007FD01) on to the instrument and obtain message code 1150 Calibrator-Set (MCC) Lot (50007FD01) expired, the actual expiration date for this lot is 2/15/2020. The customer is on ALINITY System Software v2.6.2. | A malfunction was identified with the Alinity ci System Software v2.6.2. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency of the software was not identified. PRS ICQ100233 further documents the investigation of this issue. | Yes | No
57 | The complaint text states an image of a cuvette alignment tool shown in the video associated with Maintenance and Diagnostic (M&D) procedures 4102, 4103 and 4104 does not match the actual tool used to perform the procedures. The issue occurred on Alinity SCM module SCM02543 operating at Alinity ci series System Software v2.6.2. | A review of the ALINITY product labeling and M&D procedures 4102, 4103 and 4104 indicate the images used in M&D procedure videos are for demonstration purposes only. Based on the results of the evaluation, a malfunction did not occur and there is no indication of a deficiency of the ALINITY System Software v2.6.2. Based on the results of the evaluation, the ALINITY software is performing as designed. | No | No
58 | Customer’s Issue: Out of range low results for positive control and falsely decreased values for a patient sample when using one kit of ARCHITECT ANTI-HBC II lot number 95149LI00. | A malfunction regarding out of range low result for positive control was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.The tracking and trending report review determined that there are no related adverse or non-statistical trends for the complaint issue. Based on these data, the product is performing as intended and no product issues were identified. No further investigation required. | No | No
59 | The customer observed empty conjugate bottle (error code 3000 and 3350), while using ARCHITECT BNP reagent lot 90107M800. | A deficiency regarding empty conjugate bottle (error code 3000 and 3350)was not identified as the threshold report review determined that complaint activity was not increased for the customer's issue. No additional issues were identified; no further action is required at this time. | No | No
60 | The customer observed Technopath Control level 1 out of package insert range low with the ARCHITECT Alanine Aminotransferase (ALT) assay, when using Technopath Multichem S Plus (Assayed) control lot 18609181. | The investigation by the TPM encompassed a review of batch records, quality control release data, stability data, and peer data for the ARCHITECT Alanine Aminotransferase (ALT) assay and found no issues. Furthermore, a review of CAPA and non-conformances did not identify any issues. As a result, this indicates that the product is performing as intended. A deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
61 | The complaint text states that the customer is getting error 3424 (R1 Pipettor Message Mechanism) pipettor aspiration error on the first use of reagent kits. The issue occurred on Alinity i instrument AI01727 at Alinity ci System Software v2.5.1. | A malfunction and a deficiency were both identified for the Alinity i processing module. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. PRS ICQ010174 and QR 689292 further document the investigation regarding the First reagent aspiration error out of a reagent kit from customers who are at high altitude. | Yes | No
62 | The customer observed a reagent bottle without a cap, while using ARCHITECT HCV Ag reagent kit lot 90301LP50 . | A malfunction regarding a reagent without cap was identified as the photograph and complaint text provided sufficient evidence the reagent bottle did not have a cap.A deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required at this time. | No | No
63 | The customer reported a barcode scan issue for the microparticle bottle when using ARCHITECT CMV IgM Reagent, lot 91171FN00. | A malfunction regarding barcode scan issue for microparticle was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
64 | The customer reported receiving error S0E70 Automation Retrieval List Full on their a3600 Storage and Retrieval Module (SRM) after a Daylight Saving Time (DST) occurred. An Abbott Representative contacted the customer and directed them to restart the SRM PC in order to recover from the error. Logs were retrieved for further investigation. | Based on the available information, a deficiency was not identified as there are no indications of a systemic issue with the a3600 software involving the Daylight Saving Time configuration. A malfunction was identified as the device failed to meet performance specifications or otherwise perform as intended at the customer site. This issue is being tracked through Inpeco Ticket #28071 for documentation and procedure improvements. Complaint 0560560R2352794 documents the previous investigation for this issue. | Yes | No
65 | The complaint text states that the onboard vial rack was frozen/stuck on the reagent positioner. The instrument parked onboard vial rack U7314 on the reagent positioner of the processing module and left it there for no reason until the customer manually stopped the processing module. The stuck rack was removed manually, and the instrument was reinitialized to resolve the issue. The customer is on ALINITY System Software v2.6. | A malfunction was identified for the Alinity ci System Software 2.6. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency of the software was not identified. PRS ICQ009419 further documents the investigation of this issue. | Yes | No
66 | The customer observed message code 5666 (RV loader linear queue error) and message code 5221 (motor error on Board) due to defective Reaction Vessel (RV) , when using Alinity Reaction Vessels lot 000176348. | A malfunction was identified as the description in the complaint text and attached photographs were sufficient to determine that the customer observed The customer observed message code 5666 (RV loader linear queue error) and message code 5221 (motor error on Board) due to defective Reaction Vessel (RV). A deficiency was not identified as the trend review and lot search determined that complaint activity was normal for the complaint issue. No additional issues were identified; no further action is required at this time. | No | No
67 | The customer reported observing shift up in Potassium Quality Controls results when using ICT serum Calibrator (LN 1E46-03) lot 37050UN19 on ARCHITECT c1601610. Customer states since using calibrator lot 37050UN19 level 1 has shifted up to 2.65 mmol/L when expected value was 2.58 mmol/L. | A malfunction was identified regarding shift up in Potassium Quality Controls result when using ICT serum Calibrator. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A deficiency could not be determined as the trend review and lot search did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
68 | The customer observed error code 5900 due to Reaction Vessels (RVs) with elongated gate vestiges, while using ARCHITECT Reaction Vessels lot 301662019 | A malfunction was identified regarding error code 5900 due to Reaction Vessels (RVs) with elongated gate vestiges, while using ARCHITECT Reaction Vessels. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. The complaint text and attached picture provided sufficient evidence that customer observed Reaction vessels with elongated gate vestiges causing error code 5900. A deficiency was not identified as the threshold report review determined that complaint activity was normal for the customer issue. No additional issues were identified; no further action is required at this time. | No | No
69 | The customer stated that they observed a black/ brown material in the microparticle bottle when using ARCHITECT Free T3 reagent Lot 02260UI00 | A malfunction was identified regarding black/ brown material in the microparticle bottle when using ARCHITECT Free T3 reagent. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
70 | The customer reported Non Abbott Controls generated low results out of range for Alanine Aminotransferase (ALT) reagent Lot 51144UQ05 on their ARCHITHECT analyzer. | A deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for current issue. However, a malfunction was identified. Based on the information within the complaint record, the device Alanine Aminotransferase (ALT) Reagent Lot 51144UQ05 failed to meet performance specifications or otherwise perform as intended at the customer site. No additional issues were identified; no further action is required at this time. | No | No
71 | The customer observed that the barcode could not be read, while using Alinity I Free T4 Calibrator Lot number 01050UI00. | A malfunction was identified regarding barcode could not be read, while using Alinity I Free T4 Calibrator. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A product deficiency has been identified. | Yes | No
72 | The customer observed Error code 3700, Unable to process test, (Wash Zone Aspiration 1) wash aspiration error for probe(s) (1) on ARCHITECT i2000SR, Serial Number iSR55696. | The issue in the current complaint was previously evaluated under Complaint 2252250R3787051 (see Related Evaluations). An adverse trend associated with Error 3700 was identified for the ARC WZ Temperature Tubing/Sensor 08C94-90. A deficiency has not been identified at this time. However, potential non-conformity QR 626346 has been opened to further investigate the issue | Yes | No
73 | The customer observed a software shutdown on the Alinity ci series System Software v2.6.1. Power was cycled to the SCC restoring functionality. | A malfunction was identified with the Alinity ci System Software v2.6. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency with the Alinity ci System Software v2.6.1 was not identified as the cause of the shutdown was due to an AbbottLink issue filling up the hard drive, which led to the Alinity ci system software shutting down. | No | No
74 | The customer reported imprecision for one patient sample when running Alinity c Albumin BCG LN 8P02/Lot Number 38898UN19 on Alinity c S/N AC01838. | A malfunction was identified regarding imprecision for one patient sample when running Alinity c Albumin BCG LN 8P02. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A deficiency could not be determined as the trend reviews did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
75 | A customer reported receiving Error Code 1455 with Quantia RF reagent lot 01518H000. | A malfunction and deficiency have been identified regarding Error Code 1455 with Quantia RF reagent. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site and was confirmed during testing by the third party manufacturer. QR 665920 was opened due to an Adverse complaint Trend (AssayTT-12217) to further investigate. A Field Action was issued as a result of this investigation (FA 11APR2019) and a Product Recall letter/Field Safety Notice has been issued to impacted customers, instructing them to discard lot 01518H000. | Yes | No
76 | The customer reported an increase in initial reactive results when using Alinity i HBsAg Qualitative II reagent, lot 95041FN00. They note that the reactive results are occurring when there are less than 100 tests remaining in the kit. Replacing with a new kit of the same lot resolves the issue arises again when less than 100 tests remain in the kit. | A malfunction was not identified regarding increase in initial reactive results when using Alinity i HBsAg Qualitative II reagent. Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. A systemic issue and/or product deficiency was not identified. In this case, when the samples were retested using a new reagent kit, they were negative. The HBsAg assay is a screening parameter where it is known that initial reactives happen. Instructions are provided in the product package insert regarding initial/repeat reactives. 8P10 is a highly sensitive assay which may be susceptible to interfering substances that might trigger positive assay readings. Neutralisation will reveal if the signal was generated by HBsAg or interfering substances. Lot 95041FN00 is outside the lower SPC phase limit for the positive control. Further investigation relating to a shift in controls will be performed through pNCR 682215. | Yes | No
77 | The customer observed falsely decreased Sodium results while using Alinity ICT Serum Calibrator LN 8P69-01, Lot 29084UN18 on their Alinity c analyzer. | A malfunction was identified regarding falsely decreased Sodium results while using Alinity ICT Serum Calibrator. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency was not identified as the trend review did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
78 | The complaint text states ARCHITECT Concentrated Wash Buffer for Lot Number 95465FN00 was received with crystallization around the cap. | Deficiency has been identified for crystallization on lot 95465FN00 assigned to QR639913. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise preform as intended at the customer site. | No | No
79 | The customer observed Abbott positive control out of range low and falsely elevated results for two patients, when using ARCHITECT Anti-TPO reagent lot 14092UN19. Additional Retests produced normal results. | A malfunction was identified regarding Abbott positive control out of range low and falsely elevated results for two patients, when using ARCHITECT Anti-TPO reagent. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A deficiency was not identified as the trend review and lot search determined that complaint activity was normal for the customer's issue. No additional issues were identified; no further action is required at this time. | No | No
80 | The customer observed a software shutdown on the Alinity ci series System Software v2.5.1. when moving from the specimen to control tab on the Order screen. Power was cycled to the SCC restoring functionality. | A malfunction was identified with the Alinity ci System Software v2.5.1. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, there is insufficient information to indicate a deficiency with the Alinity ci System Software v2.5.1. Per PRS ICQ008929, the issue will be addressed in a future software release. | No | No
81 | The customer called to report that Quality Control (QC) results shifted down for Total Bilirubin when using Bilirubin Calibrator LN 1E66 lot 05647FD01. The customer confirmed they used Bilirubin lot 85846FD01 and both QC levels went back to mean all in range. | A malfunction was identified as the complaint text indicated the controls were out of range. A deficiency was not identified as the threshold report review determined that this complaint is within the complaint per day rate established for 1E66_CABR. No additional issues were identified; no further action is required at this time. | No | No
82 | The customer called to report that for the Microalbumin reagent LN 2K98 lot 78868UN18 the reagent amount used didn’t match. The customer put in the new reagent, R1 detected 223 tests (instead of the 250 it had in it) and R2 detected 252 tests. The customer confirmed that R1 had fewer tests to process than R2. | The cause of the customer R1 wedge having less volume is not determined. Per the Microalbumin package insert, the estimated tests per kit is 500. Since one kit contains 2 pairs of R1 / R2 cartridges, the estimated number of tests per pair of cartridges is ~250 tests. The tests per kit may vary by individual laboratory and differences between instruments. A deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
83 | Customer's Issue: One box of ABBOTT PRISM Activator Diluent, List Number 01A75-01, lot number 01088N100 was damaged upon receipt (damage to the box only). No damage to the product bottles inside or leakage. | A malfunction (Distribution Issue) was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site (damaged product was demonstrated through plausible description and photographs). However, a product deficiency was not identified, as the 12-month review conducted for all available ADD Complaints found in the DMGD issue bucket and a review of distribution trending data identified normal activity (no systemic issue found). No additional issues were identified; no further action is required at this time. | No | No
84 | The customer reported a shift upwards for third party negative control when using Alinity i HBsAg Qualitative II reagent, lot 95041FN00. The Viroclear negative control shifted to 0.50 – 0.65 S/CO with 337 tests remaining in the reagent kit. The issue was resolved by changing to a new kit of the same lot | A malfunction regarding shift upwards for third party negative control was not identified. Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. A systemic issue and/or product deficiency was not identified. In this case, the Viroclear negative control shifted upwards to 0.50 – 0.65 S/CO but did not go out of range. | Yes | No
85 | The customer reported a First Use Failure (FUF) of list number 09D28-03 ICT Module on ARCHITECT C803422. The customer installed a new ICT Module and the instrument generated low potassium QC results. The issue was resolved by installing a different ICT Module. The ICT Module, list number 09D28-03, failed on first use, and is the subject of this evaluation. | Review of the QIST December 2019 (January 2019 – December 2019) reporting period and the ensuing search in iQ of similar complaints (on or after January 01, 2020) did not identify an adverse trend of FUF complaints involving list number 09D28-03 ICT Module. The complaint information indicates a malfunction of the part occurred; however, no systemic issue or product deficiency was identified. | No | No
86 | The customer observed discrepant results for multiple patients on the Alinity i instrument compared to another method (Roche e411), when using Alere NT-proBNP for Alinity i reagent lot 10866UP00. | A malfunction was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A Device History review was performed on the likely cause lot, which did not show any non-conformances, deviations, or potential non-conformances. Furthermore, labeling was reviewed and found to adequately address the issue under review. The trend review and lot search determined that complaint activity was normal for the complaint issue. An investigation was performed by the Third Party Manufacturer (TPM), Axis Shield. This included a review of the complaint database for similar issues, a review of in house quality control release data and stability data, and a review of the customer’s information. This investigation found no issues. As a result, this indicates that the product is performing as intended. A deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
87 | The complaint text states no Label on the Microparticle bottle when using ARCHITECT CEA Reagent, Lot Number 95030FN00. | Based on our investigation, a product deficiency regarding no label on the microparticle bottle has been identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise preform as intended at the customer site. Further investigation will be performed through QR686646. | Yes | No
88 | The customer observed Technopath Control level 3 out of package insert range low with the ARCHITECT Total Bilirubin Assay, when using Technopath Multichem S Plus (Assayed) control lot 17205173. | A malfunction was identified regarding Technopath Control level 3 out of package insert range low with the ARCHITECT Total Bilirubin Assay. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.An investigation was performed by the Third Party Manufacturer (TPM), Technopath. The investigation by the TPM encompassed a review of batch records, quality control release data, stability data, and peer data for the ARCHITECT Total Bilirubin assay and found no issues. Furthermore, a review of CAPA and non-conformances did not identify any issues. As a result, this indicates that the product is performing as intended. A deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
89 | The customer reported observing low cal absorbances when using one cartridge of Carbon Dioxide (CO2) reagent (LN 3L80) lot 52783UQ04 on ARCHITECT c1601642. The cal absorbances were approximately zero for both levels of calibrator and the slope of the calibration curve was opposite of the expected slope. The calibration was repeated successfully with a different reagent cartridge of the same lot and same lot of calibrator. | A malfunction was identified regarding low cal absorbances when using one cartridge of Carbon Dioxide (CO2) reagent. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A systemic issue and/or product deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
90 | The customer observed aberrant results when using ARCHITECT AFP Reagent, lot number 91434FN00. An Initial result of 327.80ng/ml and a repeat result of 9.59ng/ml was observed. | A malfunction regarding aberrant results when using ARCHITECT AFP Reagent was not identified. Based on the information within the complaint record, the device met performance specifications and perform as intended at the customer site. Also a systemic issue and/or product deficiency was not identified. | No | No
91 | The customer called to report that they have 2 reagent kits that failed on first use. For Lipase reagent LN 7D80 lot 54734UQ04, the kit had an insufficient error when in use even though there were enough volume still available. For Tobramycin, a new ticket is opened for a separate evaluation. | A malfunction regarding 2 reagent kits that failed on first usewas identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A deficiency could not be determined as the trend reviews and lot searches did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required. | No | No
92 | The complaint text states that the ARCHITECT analyzer generated falsely elevated CA 19-9XR results for a patient who had not been diagnosed with pancreatic cancer at the time of testing. | A malfunction was identified regarding ARCHITECT analyzer generated falsely elevated CA 19-9XR. Based on the information within the complaint record, the device did not meet performance specifications or otherwise perform as intended at the customer site. Use error may have contributed to the customer’s issue as the patient results were being compared to results obtained on another method. A cross functional team (CFT) was formed (as described in PPM ADD-CAPA-030) to review the product labeling and complaint activity. The CFT confirmed that the use described in this ticket was not in alignment with our labeling and that the current product labeling was reviewed and is sufficient and no further action will be taken at this time. Therefore, a deficiency was not identified. No additional issues were identified; no further action is required at this time. | No | No
93 | The customer observed an issue when printing calibration reports where in a blank calibration report would be printed and then the report with filled information/details on the Alinity ci series System Software v2.6.2. | A malfunction was identified with the Alinity ci System Software v2.6.2 Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.Problem Reporting System (PRS), for similar issues resulted in no PRS that can be identified for this issue. No logs were attached with this complaint. There was insufficient information to further evaluate the issue.There is insufficient information to indicate a deficiency with the Alinity ci System Software v2.6.2. | No | No
94 | The complaint text states that when the Pretrigger solution is replaced the on-board stability is 24 days and not the expected 28 days. The issue occurred on Alinity SCM01325 at software version 2.5.1 | A malfunction was identified for the Alinity ci System Software v2.5.1. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency of the software was not identified. | No | No
95 | The complaint text states there is a long delay when accessing the overall list of the on-board reagents inventory information in the Reagents screen. The issue occurred on Alinity SCM module SCM02334 at System Software v2.6.2. | A malfunction was identified with the Alinity ci System Software v2.6.2. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site.A search of the Serena Business Manager Problem Reporting System (PRS), for similar issues identified Problem PRS ICQ008410, which investigates slow screen response in the Reagent Management screen. Further investigation through PRS ICQ008410 indicated that there was no software requirement failure and the issue is still being investigated. A deficiency of the Alinity ci System Software v2.6.2 was not identified. PRS ICQ008410 further documents the investigation of the issue. | Yes | No
96 | The customer observed falsely elevated results for multiple patients after switching to a new reagent kit of ARCHITECT SCC reagent lot 86111LP31. The customer also indicated that after switching to the new reagent kit, controls were found to be out of package insert range high. | A malfunction was identified regarding elevated results for multiple patients after switching to a new reagent kit of ARCHITECT SCC. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. A deficiency was not identified as the trend review and lot search did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required at this time. | No | No
97 | The customer questioned higher than expected Troponin results for one patient when testing with ARCHITECT STAT High Sensitive Troponin-I, reagent Lot 88511UI00. The patient had an acute myocardial infarction and the values were not questioned as falsely positive but higher than expected. | A malfunction regarding higher Troponin result was not identified. Based on the information within the complaint record, the device met performance specifications and performed as intended at the customer site. In addition, a systemic issue and/or product deficiency was not identified. In this case, the patient had an acutemyocardial infarction. | No | No
98 | The customer reported insufficient fill volume when using Alinity i HBSaG Reagent, lot 03030FN00. The customer’s instrument returned error codes 3037 ((Liquid not found for pipettor (R1) at reagent carousel position (13) (outer reagent bottle)). | A malfunction regarding insufficient fill volume was identified. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a systemic issue and/or product deficiency was not identified. | No | No
99 | The complaint text states that when customer re-loaded Architect iVANC reagent Lot#01118F000 that had been calibrated on 12/01/2018, they obtained a status of No Cal. Customer recalibrated the Architect iVANC reagent to resolve the issue. The ARCHITECT i2000SR is on software version 9.25. | A malfunction was identified for the ARCHITECT system software v9.25. Based on the information within the complaint record, the device failed to meet performance specifications or otherwise perform as intended at the customer site. However, a deficiency of the software was not identified. | No | No
100 | The customer observed microparticles would not re-suspend while using, ARCHITECT Tacrolimus reagent lot 85002M800. | A malfunction was identified as the complaint text provided sufficient evidence the microparticles did not re-suspend. A deficiency was not identified as the trend reviews and lot searches did not identify an increase in complaint activity for the current issue. No additional issues were identified; no further action is required at this time. | No | No